## Slide 1: UNB MRI Centre Magnetic Field plotter & Fluid Flow with Gradient of SA3 Magnet
Devin Morin
Aug 2019

## Slide 2: Projects
Magnetic Field Plotter
SA3 Fluid Flow Measurements

## Slide 3: System parts
LakeShore 460 Gaussmeter
Velmex Stepper Motor Controllers
The MatLab script to bring it all together
MatLab Script (MagFieldMeas.m)
LakeShore Gaussmeter
Velmex Stepper Motors
Serial Port 1
Serial Port 2
-Receives a measure code, then sends the Bx, By, Bz data
Limit Switches
-Send break to motors if pressed

## Slide 4: Initial Problems
Hardware problems
Recovered drives have scarce bits of code
Little documentation

## Slide 5: Final Design GUI

## Slide 6: Handheld Unilateral Magnet Flow Design
Using SA3 with constant gradient region at some angle

## Slide 7: SA3 Fluid Flow Setup
Constant Gradient of 150 gauss/cm from 2.0cm to 2.5cm above surface of magnet
GARfield
SA3

## Slide 8: Probe at 2.26MHz
-Attempted to build an inductively matched probe as it had worked for the probe on the GARfield setup.
-Lowest frequency achieved was ~8MHz. To go lower required an inductor in series with the coil, which destroyed matching entirely.
-Overcame problems with capacitevly matched probe by decreasing the Q factor to 20 to decrease ringdown.
-Changing attenuation to match the voltage peak to peak of the probe built for the GARfield magnet.
-Approximating the pulse width

## Slide 9: Approximation of Pulse Width
Power is 11 watts given 27dB attenuation on 250watt amplifier.
= 3.6us pulse width
Experimentally the pulse width is 4us, after fine tuning.

## Slide 10: Results of the Setup
Have been able to get signal with alternating peaks that indicates fluid flow.
Alternating peaks change to be a regular exponential decay when the magnet is not at an angle.
More problems to figure out
Time(us)
Magnitude (arbitrary)
200CCM Flow CPMG with SA3 at 10 degrees

## Slide 11: Thank you
I want to say thank you for all the help I’ve received while working here. Especially to Ben, Andrés, Grant, Bruce and Jace.